หัวข้อเรื่อง	การเปรียบเทียบประสิทธิภาพของอาหารเลี้ยงเชื้อชนิดแข็งและชนิดเหลวต่อการสร้างเอนโดสปอร์ของเชื้อ Geobacillus strothermophilus ATCC 7953

ผู้วิจัย 		นางสาวภัทรวดี เลี่ยมกลาง

ที่ปรึกษา	นายสุเมธ วันภูงา

หน่วยงาน	ศูนย์ทดสอบและวิจัยคุณภาพชีววัตถุสำหรับสัตว์และ มหาวิทยาลัยราชภัฎนครราชสีมา

บทคัดย่อ

Geobacillus sterothermophilus เป็นแบคทีเรียแกรมบวก ในสกุล Bacillus ซึ่งเจริญได้ที่อุณหภูมิสูง (thermophilic bacteria)  G. stearothermophilus สามารถเติบโตในอุณหภูมิตั้งแต่ 30 - 75 ℃ เชื้อ G. sterothermophilus เมื่ออยู่ในสภาพแวดล้อมที่ไม่เหมาะสม จะเปลี่ยนรูปแบบจากเซลล์เข้าสู่สภาวะเอนโดสปอร์ ทำให้สามารถทนความร้อนได้สูงขึ้น เอนโดสปอร์ของเชื้อ G. sterthermophilus จึงถูกนำมาใช้เป็นตัวชี้วัดในการทดสอบประสิทธิภาพของเครื่องนึ่งฆ่าเชื้อ(Auto clave) โดยชุดทดสอบประสิทธิภาพการฆ่าเชื้อในท้องตลาดมีราคาสูง จึงเป็นที่มาของการเตรียมชุดทดสอบใช้เองภายในหน่วยงาน การเตรียมเอนโดสปอร์นั้นอาหารเลี้ยงเชื้อแต่ละชนิดสามารถเหนี่ยวนำให้เกิดการสร้างเอนโดสปอร์ได้แตกต่างกัน จึงทำการทดสอบประสิทธิภาพของอาหารเลี้ยงเชื้อได้แก่ TPA TPB TSA TSB TA TB NA NB ESA ESB และคำนวณหาต้นทุนของการเตรียมเอนโดสปอร์ ผลจากการศึกษาพบว่า เชื้อ G. sterthermophilus อาหารเลี้ยงเชื้อสูตร ESA และ NA ที่ใช้ระยะเวลาในการเพาะเลี้ยง 1-3 วัน ก็จะพบปริมาณเอนโดสปอร์จำนวนมากโดยพบว่า อาหารเลี้ยงเชื้อสูตร NA สามารถเตรียมเชื้อได้ 4.5 x 106 spore/ลิตรอาหารเลี้ยงเชื้อ และอาหารเลี้ยงเชื้อสูตร ESA สามารถเตรียมเชื้อได้ 7.2 x 107 spore/ลิตรอาหารเลี้ยงเชื้อ  มีต้นทุนค่าใช้จ่ายในการเตรียมอาหารเลี้ยงเชื้อ NA ต่อลิตรเท่ากับ 163 บาท สามารถเตรียมชุดทดสอบทางชีวภาพได้ทั้งหมด 22 Strip ค่าเฉลี่ยต่อชุดทดสอบ 1 ชุดเท่ากับ 7.4 บาท ส่วนอาหารเลี้ยงเชื้อของอาหารชนิด ESA มีต้นทุนค่าใช้จ่ายในการเตรียมอาหารเลี้ยงเชื้อต่อลิตรเท่ากับ 158 บาท ซึ่งสามารถเตรียมชุดทดสอบทางชีวภาพได้ทั้งหมด 288 Strip ค่าเฉลี่ยต่อชุดทดสอบ 1 ชุดเท่ากับ 0.43 บาท อาหารที่เหนี่ยวนำในการสร้างเอนโดสปอร์ที่ดีสุดคือ ESA เนื่องจากสามารถเพาะเลี้ยงเชื้อได้เอนโดสปอร์ปริมาณมากทำให้สามารถลดต้นทุนในการผลิตชุดทดสอบทางชีวภาพ อาหารเลี้ยงเชื้อ ESA จึงเหมาะสมต่อการเพาะเลี้ยงเชื้อเอนโดสปอร์มากที่สุด

คำสำคัญ: G. sterothermophillus ATCC 7953  , ชุดทดสอบประสิทธิภาพเครื่องนึ่งฆ่าเชื้อ , เอนโดสปอร์,

อาหารเลี้ยงเชื้อ

Research Title	Comparison of the effectiveness of solid and liquid media on Geobacillus sterothermophilusATCC 7953 endospore production

Researcher		Miss Pattrawadee Leamklang

Research Consultants	Mr. Sumath wanpoo-nga

Organization		Veterinary Biologics Assay and Research Center

Abstract

Geobacillus stearothermophilus is a rod-shaped, Gram-positive bacterium and a member of the division Firmicutes. There is a thermophilic bacterium that can be able grow within a temperature range of 30 to 75 oC. In addition, cell could transform be an endospore at the inappropriate conditions. It is commonly used as a challenge organism for sterilization validation studies and periodic check of sterilization cycles. Moreover, the commercially test kits that used as a standardize validation, are very expensive. The objective of this study was aimed to prepare the test kit from various culture media (; TPA, TPB, TSA, TSB, TA, TB, NA, NB, ESA, and ESB) and the effectively medium was determined for the capable of induction in the G. stearothermophilus sporulation, whereas the variable cost of the medium preparation was calculated. The results showed that spore was highly increased in the ESA and NA during the cultivation of 1 to 3 days. In the ESA, it contained spore of 7.2x107 spore/L as well as spore of 4.5x106 spore/L was observed in the NA. The variable cost of ESA preparation was about one-hundred and fifty-eight baht/L that these costs can prepare a total of 288 strips of the biological test kits. The average value of the kit was 0.43 baht/set. In contrast, the variable cost of NA preparation (/L) was one-hundred and sixty-three baht/L which can prepare a total of 22 strips of the biological indicator test kits. It indicated that the average value per test set was 7.4 baht/set. From this result, it indicated that ESA has the effectively medium to be use as biological indicator for autoclaving sterilization.

Keyword :  G. sterthermophilus, Biological indicator, Endospore, Media culture